## Cover Page
Clinical Safety Hazard Log - eConsult
Programme | eConsult Hub
Sub-Prog / Project | Clinical Safety
Prog. Director |  | Status | Issued
Owner |  | Version | 1.2
Author | Andy Harrison, ETHOS Ltd. | Version Date | 2023-07-13
Clinical Safety Hazard Log - eHub
Document Purpose
This hazard log is for the safety analysis of the eConsult "eHub" as supplied as a 'product'. This assessment is compliant to the DCB0129: Clinical Risk Management: its Application in the Manufacture of Health IT Systems

## Hazard Log
|  | Hazard Assessment |  |  |  | Initial Risk | [9 empty] | Residual Risk | [9 empty] | Owner | Status
No. |  | Hazard Description |  |  |  | Existing Controls | [5 empty] | Initial Risk Assessment |  |  |  | Additional Controls | [5 empty] | Residual Risk Assessment
[6 empty] | HIT Design |  | User Training |  | Business Process | [5 empty] | HIT Design |  | User Training |  | Business Process Change
| Category HF - Human Factors BP- Business Process Tech - Technical | Effect | Hazard | Harm | Possible Causes | Description | Evidence | Description | Evidence | Description | Evidence | Severity | Likelihood | Risk | Justification | Description | Evidence | Description | Evidence | Description | Evidence | Severity | Likelihood | Risk | Justification
eHUB 001 | BP/HF | The eHub team processing of a submitted consultation is outside of the expected time frame for the symptoms or circumstance for the consultation. | Timeframe not adhered to by Hub central process. | Potential delay in care. | The eHub team processing of a submitted consultation is outside of the expected time frame. This could be due to: 1) eConsults are left within smart inbox Waiting queue and have not been actioned 2) Insufficient staff capacity to process eConsults in timely manner 3) User error resulting in the patient not being responded to. | eHub has been thoroughly tested for all functionality including patient record update, signpost/triage, status updates, messaging and assigning to practices. | Test report / processes | eHub staff are existing experienced clinicians and administrators. Training provided including provision of an Outline PCN SoP for initial onboarding For engaged practices, baseline training provided including provision of an Outline PCN SoP for initial onboarding. | Baseline training Reference 5 | Timeframes advised and communicated at point of practice engagement with eHub and alignment confirmed. The eHub engagement with practices will be using established eHub processes which determine timescales expected for the consultation process steps. There are also established escalation processes supporting and issues which occur in processing consultations, such as inability to find a patient record, technical issues, etc. These will be reviewed on engagement by an organisation as part of an organised program of set up activities. | eHub procedures references required. | Significant | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT. The Likelihood is determined as LOW as triage times are monitored by the eHub. | System is fully tested on set-up with a new practice to ensure no issues. | [Ref 12] | Users are provided with baseline training including an Outline PCN SoP for initial onboarding of a practice. It is expected that this training material should be used to inform changes to any 'local' training at the practice. | RECOMMENDATION: Requirement for review of local training in deployment process. Practice outline SoP [Ref 2] | Timeframes advised and communicated at point of onboarding and alignment confirmed. Business Continuity Plan for resource / Technical issues | It is expected that local processes would be updated at a practice with the eHUB expected response times to be used to inform if there was a delay over their expected response times. Onboarding provides SoPs, Practice guides and issue escalation details. [Ref 4 to 8] | Significant | Very Low | 1 | SIGNIFICANT is unchanged, The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The hazard is assessed as acceptable, but it is recommended that the additional mitigation controls for the practice are considered to further reduce the likelihood and risk rating to the 'Residual' levels.
eHUB 002 | HF | The eHub team processing of a submitted consultation urgency is incorrect. | Clinical Triage incorrect | Potential delay in care. | The eHub team processing of a submitted consultation urgency is incorrect. This could be due to: - New eHub staff training. - Issue with urgency functionality. - Misinterpretation by clinician. Note, risk is with a lower urgency being assigned. Higher urgency is a resourcing risk. The following linked causes related to this hazard are covered in the existing eConsult CSCR (Ref 3). - Error in question set - eConsult Hazard - Patient expectation is for a higher urgency. | eHub has been thoroughly testing for all functionality including signpost/triage, urgency and status updates, messaging and assigning to practices. The eHub is an active service with monitoring of response times. | Test report. Process / SoP reference. | Clinicians /administrator/ receptionists have training available modules on how to review the eConsult forms urgency. For engaged practices, baseline training provided including provision of an Outline PCN SoP for initial onboarding. | Baseline training Reference 5 | For alignment of Urgency definitions, MI HUB scope of practice is provided. MI HUB inclusion criteria May 23 Existing urgency periods are predetermined and communicated. - For same day use 24 hours - Next day use 48 hours - Soon use 3 days - Routine use 2 weeks | Ref 6 | Significant | Medium | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT. Likelihood is determined as MEDIUM using the existing (initial) hazard controls. | System is fully tested for onboarding. It is expected that this will include practices setting urgencies and completing the consultation and will be familiar with the urgency functionality. | [Ref 12] | Users are provided with baseline training including an Outline PCN SoP for initial onboarding. It is expected that this training material should be used to inform changes to any 'local' training at the practice including timescales expected from urgency status.. | RECOMMENDATION: Requirement for review of local training in deployment process. | Response times for urgency are communicated at point of onboarding and alignment with practice understanding confirmed. Recommendation that the PCN to ensure there is a process in place for ensuring there is audit of clinical decisions made in eHub to ensure the quality of the interpretation of the consultation is acceptable. | It is expected that any error in urgency would be escalated by the practice for review with the eHUB team. (Part of practice incident management process). | Significant | Very Low | 1 | SIGNIFICANT is unchanged, The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The hazard is assessed as acceptable, but it is recommended that the additional mitigation controls for the practice / eHub are considered to further reduce the likelihood and risk rating to the 'Residual' level.
eHUB 003 | HF | The consultation could be marked as Closed before a key activity is performed such as booking an appointment. | Status error - status set to COMPLETE before triage decision. | Potential delay in care. | Human factors causes: - User error. - Training or lack of understanding of Status. | eHub reporting functionality can show 'completed' records with missing or incomplete information. Full audit trail available for all consultations. eConsult email to patient is automated an advises a response time on the confirmation which includes the practice contact details. Add at time of assessment 28,000 cases have passed through and no known occurrence of this hazard has been reported. | System functional Specification. | eHub staff are existing experienced clinicians and administrators. Training is provided for new administrator/ receptionists including review of the eConsult forms and setting status. | Reference 5 | There is a predetermined sequence for completion of the eConsult consultation. It is considered unlikely that a user would set an incomplete record to "Completed" status. | eHub procedures references required. | Significant | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT. Likelihood is determined as MEDIUM using the existing (initial) hazard controls. | N/a |  | Training reviews for the existing eHub staff. Training is provided for new administrator/ receptionists including review of the eConsult forms and setting status. | eHub training - RECOMMENDATION: Requirement for review of local training in deployment process. | eHub can support audit processes for checking COMPLETED records if this hazard occurs. | System functionality. | Significant | Very Low | 1 | SIGNIFICANT is unchanged. The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The hazard assessment workshop identified audit functionality which could reduce this hazard Likelihood further to LOW if formal process could be verified.
eHUB 004 | Tech | Patient registration can not be completed from the eConsult consultation. | PCN Hub unable to identify the patient or patient record. | Potential delay in care. | Patient details submitted on eConsult are unable to be registered. Patient may not be registered at any of the constituent practices. Potential for the patient to 'click-through' the practice confirmation check when completing eConsult Practices have similar sounding names and patient enters wrong practice website as a result | The eConsult patient consultation process will ask to confirm the patient practice is correct. It is unlikely that this data will be changed through the process of submission of the consultation. | System functional test. | eHub staff are existing experienced clinicians and administrators who are familiar with key data to be used when identifying patients. |  | Verification sequence: - eHub will query internal databases to locate patient. - eHub will query practice databases to locate patient. - eHub will query PDS to locate patient and identify an error in practice details. - Escalation process to highlight potential error in patient practice details (PCN). | System functional design and eConsult clinical safety report 3.3 [Ref 3] | Significant | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT. Likelihood is determined as LOW as the eConsultation will most likely be against an existing practice patient. | System is fully tested on set-up with a new practice to ensure no issues. | [Ref 12] | eHub patient identification issues are supported by verification / escalation process which should be part of onboarding. | RECOMMENDATION: Requirement for review of local training in deployment process specific to patient search process. | If the patient is unknown and cannot be matched, there is an eHub process in place to manage escalation of the issue. It is probable that the patient will be contacted as part of the process if a match cannot be made. It is expected that dependant on the patient's clinical presentation, care will be given rather than signposting to their own practice or alternative service. | eHub escalation process details provided. [Ref 4] Additionally: It is expected that practices using eConsult will have correctly registered patients. | Significant | Low | 2 | SIGNIFICANT is unchanged. The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. Hazard Likelihood further to LOW if formal eHub escalation process for could be verified.
eHUB 005 | BP | Patient may submit a form expecting an eConsultation to be processed at an earlier time. | eHub miss alignment with eConsult submission time. | Potential delay in care. | Misalignment of the availability of the eHub team and a practice communicated availability for patients to use the eConsult service. | N/A |  | eHub staff are existing experienced clinicians and administrators who are familiar with key data to be used when identifying patients. | eHub Training | Availability of the eHub team is communicated for practices to align their eConsult service. It is expected that availability of eHub staff will be inclusive of practice response times. Default time scales advised to the patients or those by the practice advised timescales are independent of eHub but is expected to be inside of eHub working times. | eHub procedures references required. | Minor | Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as MINOR as there may be an unexpected shot delay in processing the eConsult form. Likelihood is determined as LOW as the eHub availability times are communicated to practices. | N/a |  | N/A |  | Service availability times are communicated at point of onboarding and alignment with practice understanding confirmed. Recommendation: PCN would still need to ensure that the eHub availability is suitably aligned with their practice availability expectations. | It is expected that any issue with availability of service would be escalated by the practice for review with the eHUB team. [Ref 4] | Minor | Very Low | 1 | MINOR is unchanged. The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. Hazard Likelihood further to Very LOW if the practice confirms they have aligned their processes to the service availability.
eHUB 006 | BP / Tech / HF | Patient registration in the eHub can not be completed form the eConsult consultation. | Unable to assess submitted patient consultation data or consultation adequately through eConsult hub. | Potential delay in care. | eConsultation is corrupt or part of data is missing. Information in the form (e.g. free text) is contradictory to questions submitted / Contradictory to what is seen in the patient record. Human error in accessing the form. Hazard eHub 012 covers the following possible causes: Connectivity issue and Hardware errors | eHub has been thoroughly testing for all functionality including the registration and access / updates to patient records. Dual delivery of MESH to EMIS Workflow and Smart inbox would provide redundancy for current configuration. Experience indicates likelihood of a corrupt message is highly unlikely. | System functional test. Dual message configuration. | eHub staff are existing experienced clinicians and administrators who are familiar with reviewing patient consultation data. | eHub Training | It is expected that an organisation would have an exiting process for managing erroneous records and those which have clinical contradiction or data missing would highlight the issue. | Existing organisations procedure. | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as LOW as the eHub is accepting a eConsult form with a predetermined data structure. | System is fully tested on set-up with new PCN / practices to ensure eConsult forms can be submitted correctly. | [Ref 12] | Training reviews for the existing eHub staff. Training is provided for new administrator/ receptionists including review of the eConsult forms. | RECOMMENDATION: Requirement for review of local training in deployment process. | If the patients eConsult form cannot be processed, there is an eHub process in place to manage escalation of the issue. It is probable that the patient will be contacted as part of the process if a match cannot be made. | Mandatory onboarding process prior to go-live would ensure system is exercised fully end-to-end with suitable test data. eHub escalation process? Additionally: It is expected that practices using eConsult will have correct process in place for patient searches. | Significant | Very Low | 1 | SIGNIFICANT is unchanged. The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. Hazard Likelihood further to Very LOW if the practice confirms they have conducted testing of forms submission through the eHub.
eHUB 007 | Tech/ HF | The eHub and practice may be unaware of the eConsult submission. | eHub does not process or assign the eConsult report for the practice to process. | Potential delay in care. | For this to occur for the practice, the following combination of events would need to happen: Smart Inbox fails to update with inbound eConsults AND dual delivery into hub MESH inbox fails AND backup NHS email notification fails. | Because of the different mechanisms which would be required to fail simultaneously, the probability of this would be extraordinary. Current eHub configuration is redundant: For failure of the Smart Inbox, MESH messaging would be available. For failure of MESH, Smart inbox would be available. For MESH delivery and Smart Inbox delivery, the backup system is delivery via NHS Mail should those system fails. | System functional test. | N/A |  | Exiting process for business continuity for MESH and Smart Inbox | See Smart Inbox CSCR and NHS MESH | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as VERY LOW as the coincident probability of multiple failures is extremely unlikely. | Hub updates, practice messaging and email is fully tested on set-up with a new practice. | [Ref 12] | eHub technical issues are supported by escalation contacts which staff are aware of. | RECOMMENDATION: Requirement for review of local training in deployment process. | Existing Business continuity process. Recommendation for PCN / eHub to ensure NHS mail is fully supporting the failure state and that the practices are aware of the redundancy process. This includes ensuring that NHS Mail logins / passwords are kept up-to-date and managed by authorised personnel. PCN eHub SOP need to articulate the appropriate and up-to-date eConsult Support numbers to escalate technical issues. | eHub escalation process and Business continuity. Onboarding provides SoPs, Practice guides and issue escalation details. [Ref 4 to 8] | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as VERY LOW as the coincident probability of multiple failures is extremely unlikely.
eHUB 008 | Tech | eConsult form may not be able to be processed requiring the patient to be contacted. | Corruption or Loss of data within the eHub | Potential delay in care. | Functional failure in the eHub service. - System failure - Issue introduced when form is processed. - Hardware issue. | eHub uses Industry standard mechanisms to generate, message and validate data. All practices use a common platform (EMIS Web & EMIS Clinical Service in eHub) which minimises variance of data. There is an established DevOps team who monitor the service. The workflow used to process eConsults in eHub is well established. | Functional design Process for escalation of issues with form provided to PCN / eHub / Practices [Ref 4] | N/A |  | Recommendation for PCN to review incident management processes for implementation of eHub. For existing eConsult users: A process should be ensured is in place for managing eConsult forms which the pateint needs to be contacted or issues with associated records. | Existing organisational controls. | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as VERY LOW as the system design is well established and reliable. | System is fully tested on set-up with a new practice. | [Ref 12] | eHub technical issues are supported by escalation contacts which staff are aware of. | eHub training should cover points of escalation. | Existing Business continuity process reviewed on onboarding.. | eHub escalation process and Business continuity. Onboarding provides SoPs, Practice guides and issue escalation details. [Ref 4 to 8] | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as VERY LOW as the eHub process is established and can evidence a low incident of data corruption.
eHUB 009 | Tech/ HF (Resource) | All eConsult forms may not be able to be processed for a period of time. | Individuals cannot access or process eConsult submissions due to eHub service unavailability. | Potential delay in care. | eHub system or functional failure. Staff unavilable to manage the eHub Access / security issue. Failure of associated system eHub is reliant on. For access issues related to eConsult - Econsult Clinical Safety Case | The current eHub service is an established PCN resource with proven availability. The service reliability is also supported by associations with EMIS Clinical Services and the eConsult Smart Inbox (see Ref 9, 10 & 11) (Include availability of Smart inbox as mitigation for tech & resource) | Ref 9, 10 & 11 | N/A |  | Exiting business continuity process for managing system access issues. Existing business continuity plan for notifying practices to process their own eConsults within the Smart Inbox system Escalation to technical help services for service unavailability. | eConsult tech support: In app link and NHS mail tech support: https://accentureplc.service-now.com/clientportal?account=NHSmail2 | Considerable | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as CONSIDERABLE as there may be a number of patients affected. Likelihood is determined as LOW as the eHub availability is established as being reliable. | System is fully tested on set-up with a new practice. | [Ref 12] | eHub technical issues are supported by escalation contacts which staff are aware of. | eHub training should cover points of escalation. | Existing Business continuity process reviewed on onboarding.. | eHub escalation process and Business continuity. Onboarding provides SoPs, Practice guides and issue escalation details. [Ref 4 to 8] | Considerable | Very Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as CONSIDERABLE as there may be a number of patients affected. Likelihood is determined as Very LOW as the eHub availability is tested for practice onboarding.
eHUB 010 | HF / BP | Record will not be able to be viewed by practice clinician. | Patient has opted out of 3rd party sharing. | Potential delay in care. | Patient has opted out of 3rd party sharing and as such, their record will not show clinical data to any organisation outside of the patient GP Practice. | N/A - eHub functionality bound by 'opt out' protocol. | N/A | N/A |  | eHub Clinician will need to manually log into patient practice. Recommendation that eHub BP is reviewed to ensure there is an appropriate process for managing receipt of 'opt-out' patient consultations. |  | Significant | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be a number of patients affected. Likelihood is determined as LOW based on the frequency of encountering 'opt-out' patient records. | N/a |  | N/A |  | Recommendation: PCN / eHub BP to be reviewed to ensure eHub Clinician can manually log into patient practice. | PCN Recommendation | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be a number of patients affected. Likelihood is determined as VERY LOW based on the frequency of difficulty in resolving encounters with 'opt-out' patient records provided the PCN / eHub business process is reviewed for this occurrence.
Additional hazards considered in the workshop assigned as CLOSED These hazards are mitigated within the scope of assessment of eConsult or EMIS Web CSCR.
CLOSED eConsult Hazard |  |  | GP or practice clinician unable to identify the patient. |  | Hazard is mitigated as part of the eConsult product safety case. The Hub would not add additional information to the consultation outside of triage data. Ref: eConsult Safety Case Report v2.9.docx
CLOSED eConsult Hazard |  |  | A third party deliberately enters patient details through eConsult Hub in an attempt to gain prescription or some other healthcare action. |  | Hazard is mitigated as part of the eConsult product safety case. The Hub would not add additional information to the consultation outside of triage data. "Free text box leads to contradiction " - Ref: eConsult Safety Case Report v2.9.docx
CLOSED EMIS CSCR |  |  | Review of potential issues with interface with EMIS. EMIS write back CSCR supports
CLOSED | BP | N/A | Unregistered patients may submit eConsult to a GP practice. | N/A | Hazard discounted as not feasible - Patient needs to access through their registered GP or they get the advisory and details to contact the Practice. "Only patients registered with this practice can use this service" https://econsult.net/nhs-patients
CLOSED - eConsult Hazard | BP |  | Incorrect patient interacts with eConsult. |  | Patient uses incorrect identifier such as NHS number. Duplicate patient record. | Development of NHS Login functionality and NHS App utilisation mitigates this risk | eConsult clinical safety report 3.8
CLOSED - eConsult Hazard | HF |  | Free text box leads to contradiction. |  | Hazard is mitigated as part of the eConsult product safety case. The Hub would not add additional information to the consultation outside of triage data. "Free text box leads to contradiction " - Ref: eConsult Safety Case Report v2.9.docx
**Remove, this is a safety control** In this case NHS mail is a safety control for messaging. |  | Consultation details may not be received by practice. | NHS Mail fails to send/receive message | Potential delay in care. | Functional failure in the eHub or a network outage. 1- System failure 2- Network error Hazard is mitigated in part of the NHS product safety case. For failure due to NHS mail outage: Ref: https://support.nhs.net/knowledge-base/o365-shared-tenant-clinical-safety-case-report/ | eHub has ‘dual delivery’ of eConsultation requests directly into the service MESH Inbox and also the eConsult Smart Inbox Tool. There would be a rejection or system alert to warn of NHS mail failed delivery. | Functional design | N/A |  | Exiting process for managing rejected emails. Escalation for NHS Mail service unavailability. | eHub procedures references required. | Significant | Low | 2 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as LOW as the system design is reliable with 'dual messaging'. | eHub practice messaging and email is fully tested on set-up with a new practice. | [Ref 12] | eHub technical issues are supported by escalation contacts which staff are aware of. | eHub training should cover points of escalation. | Existing Business continuity process reviewed on onboarding.. | eHub escalation process and Business continuity. Onboarding provides SoPs, Practice guides and issue escalation details. [Ref 4 to 8] | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as SIGNIFICANT as there may be an requirement to reprocess the eConsult form or contact the patient. Likelihood is determined as VERY LOW as the eHub process is not reliant on a single messaging mechanism.
HAZ 004 covers this | BP | The patient may believe the form has been submitted and be waiting on a response however due to a number of reasons the eConsult form did not complete | Consultation assigned to incorrect practice. eHub HAZ 004 | Potential delay in care. | Error with eHub workflow process. Errors with Smart Inbox are covered in the Smart Inbox CSCR [Ref 9] | Practice details are included in the eConsult form and are not changed as part of the receipt and attachment to patient records. | Functional design | eHub staff are existing experienced clinicians and administrators who are familiar with matching of patient records / consultations to practices. | eHub Training | The eConsult form will be from a patient registered at a practice. It is this practice identifier which will be on the eConsult form and match the patient record practice identifier. Mitigated by patient verification sequence: - eHub will query internal databases to locate patient. - eHub will query practice databases to locate patient. - eHub will query PDS to locate patient and identify an error in practice details. - Escalation process to highlight potential error in patient practice details (PCN). | eConsult clinical safety report [Ref 2] | Significant | Very Low | 1 | The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. The SEVERITY associated with this hazard is determined as MINOR as there may be an unexpected shot delay in processing the eConsult form. Likelihood is determined as LOW as the eHub availability times are communicated to practices. | N/a |  | N/A |  | N/A |  | Significant | Very Low | 1 | SIGNIFICANT is unchanged. The conditions eConsult is used for are defined as 'non-urgent' and are not immediately life-threatening. As eConsult is not intended to be used for urgent medical advice, the cases which are triaged by eHub will not cover emergencies such as severe chest pain. Hazard Likelihood is unchanged at Very LOW with no residual mitigation required.
Note on A12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Where is eHub 008? Has this been deleted. Do the numbers need to be resequenced as only 9 hazards? Reply: We removed a hazard at the time of the workshop. I think as this is the initial copy it would be sensible to renumber. There would always be a question of 'where is 008'? Amended numbering.
Note on F5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Remove Delay in registration of patient and issue with patient record/attachment as possible causes. In our eHub case, the most likely causes: 1) eConsults are left within smart inbox Waiting queue and have not been actioned 2) Insufficient staff capacity to process eConsults in timely manner 3) User error resulting in the patient not being responded to. Reply: Amendeed
Note on F8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Consider removing “econsult unable to be attached to a matching patient record” as this is a consequence of unregistered or unmatched record. Reply: Agreed - removed
Note on G5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Does this description also cover the possible causes listed in the comment box above Reply: Yes, broad comment on the technical aspects of the risk being mitigated. The causes above are Human factors / Business process covered in the other columns.
Note on H10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: What stats here? Are we talking about the 28,000 eConsults to date since go-live? Reply: Our error - we intended putting a reference to this and this is my reminder to add to the CSCR. Now included in the document so I've removed this.
Note on K12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not sure what this description means - what does rejected eConsult forms mean? Reply: Reworded.

## Risk Matrix
Likelihood | Very High | 3 | 4 | 4 | 5 | 5 | Severity Classification | Interpretation | Number of Patients Affected
| High | 2 | 3 | 3 | 4 | 5
| Medium | 2 | 2 | 3 | 3 | 4 | Catastrophic | Death | Multiple
| Low | 1 | 2 | 2 | 3 | 4 |  | Permanent life-changing incapacity and any condition for which the prognosis is death or permanent life-changing incapacity; severe injury or severe incapacity from which recovery is not expected in the short term | Multiple
| Very Low | 1 | 1 | 2 | 2 | 3
|  | Minor | Significant | Considerable | Major | Catastrophic | Major | Death | Single
|  | Severity | [5 empty] | Permanent life-changing incapacity and any condition for which the prognosis is death or permanent life-changing incapacity; severe injury or severe incapacity from which recovery is not expected in the short term | Single
| Likelihood Category | Interpretation | [5 empty] | Severe injury or severe incapacity from which recovery is expected in the short term | Multiple
| Very high | Certain or almost certain; highly likely to occur | [5 empty] | Severe psychological trauma | Multiple
| High | Not certain but very possible; reasonably expected to occur in the majority of cases |  |  |  |  | Considerable | Severe injury or severe incapacity from which recovery is expected in the short term | Single
| Medium | Possible | [5 empty] | Severe psychological trauma | Single
| Low | Could occur but in the great majority of occasions will not | [5 empty] | Minor injury or injuries from which recovery is not expected in the short term | Multiple
| Very low | Negligible or nearly negligible possibility of occurring | [5 empty] | Significant psychological trauma | Multiple
[7 empty] | Significant | Minor injury or injuries from which recovery is not expected in the short term | Single
[8 empty] | Significant psychological trauma | Single
| 5 | Unacceptable level of risk | [5 empty] | Minor injury from which recovery is expected in the short term | Multiple
| 4 | Mandatory elimination of hazard or addition of control measure to reduce risk to an acceptable level | [5 empty] | Minor psychological upset; inconvenience | Multiple
| 3 | Undesirable level of risk. Attempts should be made to eliminate the hazard or implement control measures to reduce risk to an acceptable level. Shall only be acceptable when further risk reduction is impractical |  |  |  |  | Minor | Minor injury from which recovery is expected in the short term; minor psychological upset; inconvenience; any negligible consequence | Single
| 2 | Acceptable where cost of further reduction outweighs benefits gained or where further risk reduction is impractical
| 1 | Acceptable, no further action required

## References
Reference | Description | Version
1 | Internal Workflow Diagram | unknown
2 | SOP Practice Example | 1
3 | eConsult Clinical Safety Report | 3.3
4 | e-Consult Clinical Incident Reporting Form | unknown
5 | PCN Hub: Federated Online Consultation Triage & Minor Illness Service: [20230527_FHR_PCN_Hub_Outline_SOP] | V2.01
6 | MI HUB inclusion criteria May 23
7 | PCN Hub Operating Model – Pilot Practice Overview, Oaklands Health Centre eHub Pilot | v1.02
8 | SOP Practice Example
9 | Smart Inbox Clinical Safety Case Report | V1.0 17/10/2022
10 | EMIS Web CSCR
11 | Consultation Writeback Clinical Safety Hazard Log | PRJ0011190 June 2023
12 | UAT for initial eHub Onboarding | N/A