## Sheet1
| ADB | [8 empty] | Room Data Sheet | [11 empty] | D0205
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | D0205 |  |  |  |  | Rest room with beverage, snack preparation bay: 5 staff
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Rest and relaxation or sleeping. 2) Patient, relative or staff may be reading. 3) Preparing beverages and snacks. 4) Consuming beverages, meals and snacks. 5) Holding/ storing food in refrigerator. 6) Holding crockery and cutlery. 7) Washing up of crockery and cutlery. 8) Staff notices, information and/or messages may be displayed. 9) Informal staff discussions. 10) Hand washing.
| Personnel: | [6 empty] | 5 x Staff
| Planning Relationships: | [5 empty] | Part of staff accommodation.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 11 |  |  |  |  | Height (mm): |  |  |  | 2400
| Notes: | [7 empty] | Separate data and voice outlets may be used where structure cabling solutions are not available.
| ADB | [7 empty] | Room Environmental Data
D0205
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | D0205 |  |  |  |  | Rest room with beverage, snack preparation bay: 5 staff
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 – 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 22
| Minimum Air Changes (AC/hr): | [13 empty] | 6
| Ventilation Type: | [13 empty] | S/E/N
| Pressure Relative to Adjoining Space: | [13 empty] | -ve
| Supply Air: Final Filter Class | [12 empty] | G4
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | S/V/EM
| Daytime General Service Illuminance (Lux): | [13 empty] | 50 |  | Floor
| Daytime Specific Service Illuminance (Lux): | [13 empty] | 200 |  | On tables
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
| Local Task Illuminance (Lux): | [13 empty] | 300 |  | At servery counter
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 90
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
| General Notes: | [5 empty] | Refer to SLL Lighting Guide 2 for more detailed guidance
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Moderate |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Medium
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41/>55
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | D0205
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | D0205 |  |  |  |  | Rest room with beverage, snack preparation bay: 5 staff
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 1 doorset: 1 x personnel, wheelchair & equipment access (1000mm)
| Windows: | [5 empty] | Desirable - Project Option Clear glass with solar and privacy control Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | D0205
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | D0205 |  |  |  |  | Rest room with beverage, snack preparation bay: 5 staff
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | CAT002 |  | MINI KITCHEN (combined unit) with refrigerator and microwave, sink top, wall cabinet, cupboard and drawers, laminate backboard, 1970H 1200W 600D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT052 |  | CONNECTION UNIT, switched, 13 amp | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT131 |  | SOCKET outlet data/voice, double. | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SMC073 |  | CABINET, upper, modular storage, for 600 facing inserts, with door, 750H | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SWC025 |  | SWITCH, light | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WAT003 |  | WATER BOILER, 2.5 litre, electric, wall mounted | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA013 |  | BOARD, display/notice, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS013 |  | DISPENSER, paper towel, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | RAC091 |  | RACK, magazine, 3 shelves, floor standing, fixed to the wall, 1400H 900W 300D | [16 empty] | 2
| 3 |  |  | 3 |  |  |  |  | CHA305 |  | CHAIR, easy, low back, upholstered, wipeable | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | CHA318 |  | CHAIR, upright, with arms, upholstered, stacking, wipeable | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | HOL006 |  | HOLDER, sack, with lid foot operated, medium, freestanding, 875H 430W 385D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TAB056 |  | TABLE, occasional, round, 415H 610mm dia. | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | K0912
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | K0912 |  |  |  |  | Computer communication: IT hub room
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Computer workstation(s) may be used.
| Personnel: | [6 empty] | Up to 3 staff
| Planning Relationships: | [5 empty] | Close to areas served. Easy access to General office, administration and records. Location: Maximum 80m away from the furthest PC workstation.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 9 |  |  |  |  | Height (mm): |  |  |  | 2700
| Notes: | [7 empty] | LAN - Local Area Network Equipment: Project specific. Electrical trunking to main walls. Separate data and voice outlets may be used where structure cabling solutions are not available.
| ADB | [7 empty] | Room Environmental Data
K0912
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | K0912 |  |  |  |  | Computer communication: IT hub room
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 25
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 20
| Minimum Air Changes (AC/hr): | [13 empty] | 6
| Ventilation Type: | [13 empty] | S/E/N
| Pressure Relative to Adjoining Space: | [13 empty] | Bal or -ve
| Supply Air: Final Filter Class | [12 empty] | G4
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 19
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 80
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 80 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | Typical
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke |  | Air Sampling System to be considered
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | K0912
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | K0912 |  |  |  |  | Computer communication: IT hub room
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 1 doorset: 1 x personnel & equipment access (1000mm)
| Windows: | [5 empty] | Not required
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | K0912
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | K0912 |  |  |  |  | Computer communication: IT hub room
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 12 |  |  | 12 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 12 |  |  | 12 |  |  |  |  | OUT131 |  | SOCKET outlet data/voice, double. | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SWC025 |  | SWITCH, light | [16 empty] | 1
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1021
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1021 |  |  |  |  | Wash room: 2 washer-disinfectors, single chamber No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Unloading trolleys received from Contaminated Returns Hold. 2) Checking and sorting items for processing. 3) Disassembling and preparing medical devices and equipment for processing. 4) Loading of medical devices into washer-disinfectors. 5) Loading of medical devices into ultrasonic. 6) Cleaning of medical devices and accessories in sink. 7) Transfer processed items to the IAP room. 8) Small items of equipment are held stored 9) Receiving materials, consumables and small items via hatch. 10) Holding/storing cleaning equipment. 11) Hand rinsing.
| Personnel: | [6 empty] | 2 - 3 staff
| Planning Relationships: | [5 empty] | Adjacent to IAP room. Direct access to the contaminated returns lobby. Direct access to general waste disposal/laundry returns area. Direct access to materials transfer hatches. Direct access to gowning room. Direct access to domestic services room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 54 |  |  |  |  | Height (mm): |  |  |  | 2700
Manual sorting workstations 2 x 3.0 sq.m per place = 6.0 sq.m Manual washing workstation 1 x 4.5 sq.m per place = 4.5 sq.m Single chamber washer/disinfectors 2 x 6.0 sq.m per place = 12.0 sq.m Handling systems 3 x 4.0 sq.m = 12.0 sq.m Net total = 34.5 sq.m Circulation within space & support 40% = 14.0 sq.m Storage 6.0 sq.m = 6.0 sq.m Gross Total = 54.5 sq.m
| Notes: | [7 empty] | Two-way communication system should be installed between wash room and IAP room, adjacent to hatches. IT - see HBN SSD Para 4.44 Pass-through hatch to IAP room for manually cleaned items. Pass-through hatch from IAP room for return carriages and medical devices. Pass-through hatch from materials store.
| ADB | [7 empty] | Room Environmental Data
R1021
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1021 |  |  |  |  | Wash room: 2 washer-disinfectors, single chamber No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | -ve
| Supply Air: Final Filter Class | [12 empty] | F9
Permisiible Relative Humidity Range (%):
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1021
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1021 |  |  |  |  | Wash room: 2 washer-disinfectors, single chamber No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 2 sets of doors: 2 x personnel, trolley & equipment access (1700mm) The needs for door depend on the plan adopted.
| Windows: | [5 empty] | Desirable - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Essential Two pass-through hatches, 2 sliding interlocking doors, all corners of internal surfaces of hatch to be coved. One pass-through hatch from materials store Designation to be validated against current documentation.
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1021
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1021 |  |  |  |  | Wash room: 2 washer-disinfectors, single chamber No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 2 |  |  | 2 |  |  |  |  | BAS105 |  | BASIN, small, general pattern, vitreous china, 1 tap right hand hole, no overflow, bottom outlet, 400W 300D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSC187 |  | CABINET base, 400mm facing, with 2 shelves, 1 door hinged right, on plinth, o/a height 900 | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSC188 |  | CABINET base, 400mm facing, with 2 shelves, 1 door hinged left, on plinth, o/a height 900 | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSW068 |  | WORKTOP, for 400mm facing inserts cabinets, 1800W 700D nominal | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT075 |  | OUTLET isolator 20amp, TP&N. | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT210 |  | SOCKET outlet two-way communication system (intercom). | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT453 |  | OUTLET, 4 kPa compressed air, medical | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | TAP289 |  | TAP, monobloc, pillar mixer, integral thermostatic, short lever. | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | TAP365 |  | *TAP pillar, mixer, double lever action, hose outlet, spray gun | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR049 |  | BENCH, stainless steel, with 2 sink bowls (280x910x600) and drainers, undershelf, o/a 900H 3000W 700D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR055 |  | WORKSTATION TABLE, with canopy, solid back and sides transparent front, high/low open, 1600H 1117W 750D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA037 |  | BOARD, marker, whiteboard, dry-wipe, with pen holder, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | CLO001 |  | CLOCK battery, wall mounted | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | DEI001 |  | DEIONISER, water, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS007 |  | DISPENSER, paper towel roll, wall mounted | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | DIS013 |  | DISPENSER, paper towel, wall mounted | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | DIS030 |  | DISPENSER, soap, disposable single cartridge, lever action, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS436 |  | DISPENSER, disposable apron, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | WAS051 |  | WASH, eye, first aid, with mirror, wall mounted, 360H 250W 100D | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | WAS056 |  | WASHER-DISINFECTOR, pass-through, single chamber, 10 DIN baskets, vertical double door, o/a 1800H 1230W 880D | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | HOL006 |  | HOLDER, sack, with lid foot operated, medium, freestanding, 875H 430W 385D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | HOL014 |  | HOLDER, sack, with lid foot operated, large, capacity 120 litre, mobile | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | RAC410 |  | RACK, metal, 1 bay, 4 shelves, 1980H 1200W 600D | [16 empty] | 3
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1021
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1021 |  |  |  |  | Wash room: 2 washer-disinfectors, single chamber No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | SPR001 |  | SPRAYGUN, water, with nozzle attachments and hose | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TRO404 |  | TROLLEY, basket/container transporting, for vertical door washer/desinfectors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | ULT006 |  | ULTRASONIC CLEANER, bench top, (tank size 150x495x280mm), Includes basket and stainless steel lid, heated, 280H 535W 355D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | WOR050 |  | WORKBENCH, stainless steel, undershelf, 900H 1600W 650D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | WOR051 |  | WORKBENCH, stainless steel, undershelf, 900H 800W 650D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1023
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1023 |  |  |  |  | Wash room: 2 washer-disinfectors, multi chamber No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Unloading trolleys received from Contaminated Returns Hold. 2) Checking and sorting items for processing. 3) Disassembling and preparing medical devices and equipment for processing. 4) Loading of medical devices into washer-disinfectors. 5) Loading of medical devices into ultrasonic. 6) Cleaning of medical devices and accessories in sink. 7) Transfer processed items to the IAP room. 8) Small items of equipment are held stored 9) Receiving materials, consumables and small items via hatch. 10) Holding/storing cleaning equipment. 11) Hand rinsing.
| Personnel: | [6 empty] | 2 - 3 staff
| Planning Relationships: | [5 empty] | Adjacent to IAP room. Direct access to the contaminated returns lobby. Direct access to general waste disposal/laundry returns area. Direct access to materials transfer hatches. Direct access to gowning room. Direct access to domestic services room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 76.5 |  |  |  |  | Height (mm): |  |  |  | 2700
Manual sorting workstations 3 x 3.0 sq.m per place = 9.0 sq.m Manual washing workstation 2 x 4.5 sq.m per place = 9.0sq.m Multi chamber washer/disinfectors 2 x 8.0 sq.m per place = 16.0 sq.m Handling systems 3 x 4.0 sq.m = 12.0 sq.m Net total = 46.0 sq.m Circulation within space & support 40% = 18.5 sq.m Storage 12.0 sq.m = 12.0 sq.m Gross Total = 76.5 sq.m
| Notes: | [7 empty] | Two-way communication system should be installed between wash room and IAP room, adjacent to hatches. IT - see HBN SSD Para 4.44 Pass-through hatch to IAP room for manually cleaned items. Pass-through hatch from IAP room for return carriages and medical devices. Pass-through hatch from materials store.
| ADB | [7 empty] | Room Environmental Data
R1023
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1023 |  |  |  |  | Wash room: 2 washer-disinfectors, multi chamber No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | -ve
| Supply Air: Final Filter Class | [12 empty] | F9
Permisiible Relative Humidity Range (%):
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1023
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1023 |  |  |  |  | Wash room: 2 washer-disinfectors, multi chamber No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 2 sets of doors: 2 x personnel, trolley & equipment access (1700mm) The needs for door depend on the plan adopted.
| Windows: | [5 empty] | Desirable - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Essential Two pass-through hatches, 2 sliding interlocking doors, all corners of internal surfaces of hatch to be coved. One pass-through hatch from materials store Designation to be validated against current documentation.
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1023
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1023 |  |  |  |  | Wash room: 2 washer-disinfectors, multi chamber No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | BAS105 |  | BASIN, small, general pattern, vitreous china, 1 tap right hand hole, no overflow, bottom outlet, 400W 300D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSC187 |  | CABINET base, 400mm facing, with 2 shelves, 1 door hinged right, on plinth, o/a height 900 | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSC188 |  | CABINET base, 400mm facing, with 2 shelves, 1 door hinged left, on plinth, o/a height 900 | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | MSW068 |  | WORKTOP, for 400mm facing inserts cabinets, 1800W 700D nominal | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT075 |  | OUTLET isolator 20amp, TP&N. | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT210 |  | SOCKET outlet two-way communication system (intercom). | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT453 |  | OUTLET, 4 kPa compressed air, medical | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | TAP289 |  | TAP, monobloc, pillar mixer, integral thermostatic, short lever. | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | TAP365 |  | *TAP pillar, mixer, double lever action, hose outlet, spray gun | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR049 |  | BENCH, stainless steel, with 2 sink bowls (280x910x600) and drainers, undershelf, o/a 900H 3000W 700D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR055 |  | WORKSTATION TABLE, with canopy, solid back and sides transparent front, high/low open, 1600H 1117W 750D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA037 |  | BOARD, marker, whiteboard, dry-wipe, with pen holder, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | CLO001 |  | CLOCK battery, wall mounted | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | DEI001 |  | DEIONISER, water, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS007 |  | DISPENSER, paper towel roll, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS013 |  | DISPENSER, paper towel, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS030 |  | DISPENSER, soap, disposable single cartridge, lever action, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS436 |  | DISPENSER, disposable apron, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | WAS051 |  | WASH, eye, first aid, with mirror, wall mounted, 360H 250W 100D | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | WAS058 |  | WASHER-DISINFECTOR, automatic, multi chamber, 3 x 10 DIN baskets, vertical double door, complete | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | HOL006 |  | HOLDER, sack, with lid foot operated, medium, freestanding, 875H 430W 385D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | HOL014 |  | HOLDER, sack, with lid foot operated, large, capacity 120 litre, mobile | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | RAC410 |  | RACK, metal, 1 bay, 4 shelves, 1980H 1200W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | SPR001 |  | SPRAYGUN, water, with nozzle attachments and hose | [16 empty] | 3
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1023
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1023 |  |  |  |  | Wash room: 2 washer-disinfectors, multi chamber No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 2 |  |  | 2 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TRO404 |  | TROLLEY, basket/container transporting, for vertical door washer/desinfectors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | ULT006 |  | ULTRASONIC CLEANER, bench top, (tank size 150x495x280mm), Includes basket and stainless steel lid, heated, 280H 535W 355D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | WOR050 |  | WORKBENCH, stainless steel, undershelf, 900H 1600W 650D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | WOR051 |  | WORKBENCH, stainless steel, undershelf, 900H 800W 650D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1032
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1032 |  |  |  |  | Wash area: 2 pass-through trolley washers No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Washing and cleaning trolleys.
Personnel:
| Planning Relationships: | [5 empty] | Adjacent to wash room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 15 |  |  |  |  | Height (mm): |  |  |  | 2400
Based upon 7.5 sq.m per trolley washer incl. access space, etc.
| Notes: | [7 empty] | Min two trolley washers required to allow for down time.
| ADB | [7 empty] | Room Environmental Data
R1032
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1032 |  |  |  |  | Wash area: 2 pass-through trolley washers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | -ve
| Supply Air: Final Filter Class | [12 empty] | F9
Permisiible Relative Humidity Range (%):
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1032
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1032 |  |  |  |  | Wash area: 2 pass-through trolley washers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Doorsets:
| Windows: | [5 empty] | Desirable - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1032
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1032 |  |  |  |  | Wash area: 2 pass-through trolley washers No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 2 |  |  | 2 |  |  |  |  | WAS059 |  | WASHER PASS THROUGH, single trolley | [16 empty] | 2
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1041
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1041 |  |  |  |  | Inspection, assembly, packing (IAP) room: 2 workstations Served 2 single washer/disinfectors No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Receiving clean and dry items from washing area. 2) Inspecting, function testing, assembling and packing cleaned and disinfected medical devices and raw materials in preparation for sterilization. 3) Inspecting and testing items requiring a light source/magnification/ medical compressed air. 4) Prepare packs and heat seal. 5) Transfer packaged goods to the packed products transfer facility. 6) Moving/manoeuvring sterilizer loading trolley. 7) Receive empty carriers/pallets via double hatch from sterilizer loading area. 8) Containers and carriers are returned to the wash room. 9) Maintenance of quality assurance/traceability records.
| Personnel: | [6 empty] | 2 - 3 Staff
| Planning Relationships: | [5 empty] | Direct access to/from gowning room. Adjacent to washing area with airlock transfer facilities. Adjacent to sterilizer loading area with airlock transfer facilities. Direct access to domestic services room. Direct access to materials transfer room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 54 |  |  |  |  | Height (mm): |  |  |  | 2700
General space for support activities, etc. = 20.0 sq.m Workstations 2 x 5.0 sq.m per place = 10.0 sq.m Single-chamber washer/disinfector exit 2 x 2.0 sq.m per place = 4.0 sq.m Handling system 3 x 2.0 sq.m = 6.0 sq.m Storage = 14.0 sq.m Gross total = 54.0 sq.m
| Notes: | [7 empty] | A demarcated area within IAP room should be provided to segregate non-conforming products. Two-way communication system should be installed between wash room and IAP room, adjacent to hatches. IT - see HBN SSD Para 4.89 Telephone, wall mounted, recessed, hand free operation, voice switching, spray cleanable. Service-maintenance access points should be planned out of this area, see HBN SSD Para 3.42 and 3.43
| ADB | [7 empty] | Room Environmental Data
R1041
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1041 |  |  |  |  | Inspection, assembly, packing (IAP) room: 2 workstations Served 2 single washer/disinfectors No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | POS
| Supply Air: Final Filter Class | [12 empty] | >H10
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1041
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1041 |  |  |  |  | Inspection, assembly, packing (IAP) room: 2 workstations Served 2 single washer/disinfectors No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 3 sets of doors: 3 x personnel, trolley & equipment access (1000mm)
| Windows: | [5 empty] | Desirable - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Essential Designation to be validated against current documentation. HTM 57 (Mar-2005)
| Hatch: | [5 empty] | Essential Two pass-through hatches, 2 sliding interlocking doors, all corners of internal surfaces of hatch to be coved. One pass-through hatch from materials store Designation to be validated against current documentation.
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1041
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1041 |  |  |  |  | Inspection, assembly, packing (IAP) room: 2 workstations Served 2 single washer/disinfectors No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 2 |  |  | 2 |  |  |  |  | RAC168 |  | *RACK, tote box, single sided, with containers, louvered panel, 1370x480, mobile, o/a1500H 480W 460D.
| 1 |  |  | 1 |  |  |  |  | CLO005 |  | CLOCK impulse, wall mounted | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT002 |  | OUTLET cable, 13 amp | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT453 |  | OUTLET, 4 kPa compressed air, medical | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | RAC420 |  | RACK, metal, 1 bay, 4 shelves, 2000H 914W 600D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | RAC422 |  | RACK, metal, 2 bays, 4 shelves, 2000H 1800W 600D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR055 |  | WORKSTATION TABLE, with canopy, solid back and sides transparent front, high/low open, 1600H 1117W 750D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA037 |  | BOARD, marker, whiteboard, dry-wipe, with pen holder, wall mounted, 900H 1200W | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | CON047 |  | CONVEYOR, roller, gravity feed, stainless steel, 850H 2000W 485D | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | LAB033 |  | BENCH MAGNIFIER, with fluorescent lamp, wall mounted | [16 empty] | 2
| 3 |  |  | 3 |  |  |  |  | CHA004 |  | CHAIR, height adjustable 540-790, tilting back, adjustable footrest, swivel, 5 star base, on castors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM032 |  | COMPUTER VDT MONITOR | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM033 |  | COMPUTER KEYBOARD | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM038 |  | COMPUTER PRINTER, laser, A4, 250H 380W 385D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | HEA010 |  | HEAT SEALING DEVICE, Includes paper guide, 150H 550W 250D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | HOL014 |  | HOLDER, sack, with lid foot operated, large, capacity 120 litre, mobile | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | RAC252 |  | RACK, wrapping paper, mobile, 1000H 1100W 400D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TAB011 |  | TABLE, 725H 1220W 750D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TAB186 |  | TABLE, with supports, 900H 1120W 750D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TRA070 |  | TRAY, guide, for heat sealing device, (Part of HEA010) | [16 empty] | 3
| 3 |  |  | 3 |  |  |  |  | TRO021 |  | TROLLEY, 4 sets of runners, 850H 600W 600D | [16 empty] | 3
| 3 |  |  | 3 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | WOR041 |  | WORKSTATION TABLE, stainless steel, height adjustable, mobile with brake, 900H 1600W 800D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1043
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1043 |  |  |  |  | Inspection, assembly, packing (IAP) room: 10 workstations Served by 2 multi chamber washer/disinfectors No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Receiving clean and dry items from washing area. 2) Inspecting, function testing, assembling and packing cleaned and disinfected medical devices and raw materials in preparation for sterilization. 3) Inspecting and testing items requiring a light source/magnification/ medical compressed air. 4) Prepare packs and heat seal. 5) Transfer packaged goods to the packed products transfer facility. 6) Moving/manoeuvring sterilizer loading trolley. 7) Receive empty carriers/pallets via double hatch from sterilizer loading area. 8) Containers and carriers are returned to the wash room. 9) Maintenance of quality assurance/traceability records.
| Personnel: | [6 empty] | 4 - 6 Staff
| Planning Relationships: | [5 empty] | Direct access to/from gowning room. Adjacent to washing area with airlock transfer facilities. Adjacent to sterilizer loading area with airlock transfer facilities. Direct access to domestic services room. Direct access to materials transfer room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 104 |  |  |  |  | Height (mm): |  |  |  | 2700
General space for support activities, etc. = 20.0 sq.m Workstations 10 x 5.0 sq.m per place = 50.0 sq.m Multi-chamber washer/disinfector exit 2 x 2.0 sq.m per place = 4.0 sq.m Handling system 3 x 2.0 sq.m = 6.0 sq.m Storage = 24.0 sq.m Gross total = 104.0 sq.m
| Notes: | [7 empty] | A demarcated area within IAP room should be provided to segregate non-conforming products. Two-way communication system should be installed between wash room and IAP room, adjacent to hatches. IT - see HBN SSD Para 4.89 Telephone, wall mounted, recessed, hand free operation, voice switching, spray cleanable. Service-maintenance access points should be planned out of this area, see HBN SSD Para 3.42 and 3.43
| ADB | [7 empty] | Room Environmental Data
R1043
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1043 |  |  |  |  | Inspection, assembly, packing (IAP) room: 10 workstations Served by 2 multi chamber washer/disinfectors No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | POS
| Supply Air: Final Filter Class | [12 empty] | >H10
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1043
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1043 |  |  |  |  | Inspection, assembly, packing (IAP) room: 10 workstations Served by 2 multi chamber washer/disinfectors No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 3 sets of doors: 3 x personnel, trolley & equipment access (1000mm)
| Windows: | [5 empty] | Desirable - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Essential Designation to be validated against current documentation. HTM 57 (Mar-2005)
| Hatch: | [5 empty] | Essential Two pass-through hatches, 2 sliding interlocking doors, all corners of internal surfaces of hatch to be coved. One pass-through hatch from materials store Designation to be validated against current documentation.
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1043
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1043 |  |  |  |  | Inspection, assembly, packing (IAP) room: 10 workstations Served by 2 multi chamber washer/disinfectors No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | CLO005 |  | CLOCK impulse, wall mounted | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT002 |  | OUTLET cable, 13 amp | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT453 |  | OUTLET, 4 kPa compressed air, medical | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | RAC420 |  | RACK, metal, 1 bay, 4 shelves, 2000H 914W 600D | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | RAC422 |  | RACK, metal, 2 bays, 4 shelves, 2000H 1800W 600D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR055 |  | WORKSTATION TABLE, with canopy, solid back and sides transparent front, high/low open, 1600H 1117W 750D | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | BOA037 |  | BOARD, marker, whiteboard, dry-wipe, with pen holder, wall mounted, 900H 1200W | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | CON048 |  | CONVEYOR, roller, gravity feed, stainless steel, 850H 2000W 680D | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | LAB033 |  | BENCH MAGNIFIER, with fluorescent lamp, wall mounted | [16 empty] | 2
| 12 |  |  | 12 |  |  |  |  | CHA004 |  | CHAIR, height adjustable 540-790, tilting back, adjustable footrest, swivel, 5 star base, on castors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM032 |  | COMPUTER VDT MONITOR | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM033 |  | COMPUTER KEYBOARD | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM038 |  | COMPUTER PRINTER, laser, A4, 250H 380W 385D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | HEA010 |  | HEAT SEALING DEVICE, Includes paper guide, 150H 550W 250D | [16 empty] | 3
| 5 |  |  | 5 |  |  |  |  | HOL014 |  | HOLDER, sack, with lid foot operated, large, capacity 120 litre, mobile | [16 empty] | 3
| 4 |  |  | 4 |  |  |  |  | RAC167 |  | RACK, tote box, single sided, with containers, louvered panel 1080x900, mobile, o/a 1195H 955W 460D | [16 empty] | 3
| 4 |  |  | 4 |  |  |  |  | RAC252 |  | RACK, wrapping paper, mobile, 1000H 1100W 400D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TAB011 |  | TABLE, 725H 1220W 750D | [16 empty] | 3
| 3 |  |  | 3 |  |  |  |  | TAB186 |  | TABLE, with supports, 900H 1120W 750D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | TRA070 |  | TRAY, guide, for heat sealing device, (Part of HEA010) | [16 empty] | 3
| 12 |  |  | 12 |  |  |  |  | TRO021 |  | TROLLEY, 4 sets of runners, 850H 600W 600D | [16 empty] | 3
| 7 |  |  | 7 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| 10 |  |  | 10 |  |  |  |  | WOR041 |  | WORKSTATION TABLE, stainless steel, height adjustable, mobile with brake, 900H 1600W 800D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1052
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1052 |  |  |  |  | Loading/unloading area: 2 sterilizers No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Receiving packaged products before sterilization. 2) Parking sterilizer loading trolleys/carriages. 3) Moving/manoeuvring trolley(s). 4) Sterilizer is loaded/unloaded. 5) Processes and tests are recorded. 6) Storing records and test charts. 7) Storing test material. 8) Labelling of sterilized packs. 9) Segregate and hold (a) items visibly defective following sterilization, (b) contents failed process cycle. 10) Maintenance of sterilizers. 11) Transfer empty carriers/pallets to packing area
| Personnel: | [6 empty] | Up to 2 persons
| Planning Relationships: | [5 empty] | Direct access to cooling area. Direct access to packaged products transfer facilities. Direct access to sterilizer plant room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 29.5 |  |  |  |  | Height (mm): |  |  |  | 2700
Area based upon 8.5 sq.m per sterilizer incl. carriage parking plus 2.5 sq.m for administration functions, etc. Test pack conditioning space - 10 sq.m for testing 1 sterilizer at any one time 20 sq.m or testing 2 sterilizer at any one time
| Notes: | [7 empty] | For IT and communications the following equipment may be needed: telephone/voice communication system; public address system; IT terminal, writing surface and hand held data code readers, see HBN SSD Para 4.131 - 4.132
| ADB | [7 empty] | Room Environmental Data
R1052
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1052 |  |  |  |  | Loading/unloading area: 2 sterilizers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | NEG
| Supply Air: Final Filter Class | [12 empty] | F9
Permisiible Relative Humidity Range (%):
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 55
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1052
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1052 |  |  |  |  | Loading/unloading area: 2 sterilizers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Not applicable (subject to plan adopted and ventilation requirements).
| Windows: | [5 empty] | Non-essential - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1052
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1052 |  |  |  |  | Loading/unloading area: 2 sterilizers No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | CLO005 |  | CLOCK impulse, wall mounted | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT002 |  | OUTLET cable, 13 amp | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | SHE241 |  | SHELF, 1800W 300D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | STF135 |  | STORAGE UNIT, lower, cupboard, 2 door, 1 shelf, 550H 1000W 450D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WKT148 |  | WORKTOP, cantilevered from wall, 1800W 650D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA013 |  | BOARD, display/notice, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | CAB025 |  | CABINET, filing, 4 drawer, 1320H 465W 620D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | CAR020 |  | CARRIER, sterilizer, internal load, standard mesh carrier with 2 shelves, 21cuft capacity | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | PAL010 |  | PALLET, sterilizer, internal load, 200H 400W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | STO003 |  | STOOL, height adjustable, on slides, 600-750H 320 dia. | [16 empty] | 3
| 4 |  |  | 4 |  |  |  |  | TRO011 |  | TROLLEY, sterilizer loading, 900H 700W 1300D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1053
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1053 |  |  |  |  | Loading/unloading area: 3 sterilizers No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Receiving packaged products before sterilization. 2) Parking sterilizer loading trolleys/carriages. 3) Moving/manoeuvring trolley(s). 4) Sterilizer is loaded/unloaded. 5) Processes and tests are recorded. 6) Storing records and test charts. 7) Storing test material. 8) Labelling of sterilized packs. 9) Segregate and hold (a) items visibly defective following sterilization, (b) contents failed process cycle. 10) Maintenance of sterilizers. 11) Transfer empty carriers/pallets to packing area
| Personnel: | [6 empty] | Up to 4 persons
| Planning Relationships: | [5 empty] | Direct access to cooling area. Direct access to packaged products transfer facilities. Direct access to sterilizer plant room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 38 |  |  |  |  | Height (mm): |  |  |  | 2700
Area based upon 8.5 sq.m per sterilizer incl. carriage parking plus 2.5 sq.m for administration functions, etc. Test pack conditioning space - 10 sq.m for testing 1 sterilizer at any one time 20 sq.m or testing 2 sterilizer at any one time
| Notes: | [7 empty] | For IT and communications the following equipment may be needed: telephone/voice communication system; public address system; IT terminal, writing surface and hand held data code readers, see HBN SSD Para 4.131 - 4.132
| ADB | [7 empty] | Room Environmental Data
R1053
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1053 |  |  |  |  | Loading/unloading area: 3 sterilizers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | NEG
| Supply Air: Final Filter Class | [12 empty] | F9
Permisiible Relative Humidity Range (%):
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 22
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 55
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1053
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1053 |  |  |  |  | Loading/unloading area: 3 sterilizers No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Not applicable (subject to plan adopted and ventilation requirements).
| Windows: | [5 empty] | Non-essential - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1053
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1053 |  |  |  |  | Loading/unloading area: 3 sterilizers No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | CLO005 |  | CLOCK impulse, wall mounted | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT002 |  | OUTLET cable, 13 amp | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT121 |  | SOCKET outlet, computer data, single | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT215 |  | SOCKET outlet, telephone | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | SHE241 |  | SHELF, 1800W 300D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | STF135 |  | STORAGE UNIT, lower, cupboard, 2 door, 1 shelf, 550H 1000W 450D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WKT148 |  | WORKTOP, cantilevered from wall, 1800W 650D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA013 |  | BOARD, display/notice, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | CAB025 |  | CABINET, filing, 4 drawer, 1320H 465W 620D | [16 empty] | 3
| 3 |  |  | 3 |  |  |  |  | CAR020 |  | CARRIER, sterilizer, internal load, standard mesh carrier with 2 shelves, 21cuft capacity | [16 empty] | 3
| 4 |  |  | 4 |  |  |  |  | PAL010 |  | PALLET, sterilizer, internal load, 200H 400W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | STO003 |  | STOOL, height adjustable, on slides, 600-750H 320 dia. | [16 empty] | 3
| 6 |  |  | 6 |  |  |  |  | TRO011 |  | TROLLEY, sterilizer loading, 900H 700W 1300D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | TRO205 |  | TROLLEY, stainless steel, 1 shelf, 900H 600W 600D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | R1080
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1080 |  |  |  |  | Test room: equipment and data SSD No room layout
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Administration 2) Computer workstation(s) may be used. 3) Compiling and maintaining records. 4) Information/data is recorded 5) Storing test equipment. 6) Records are stored. 7) Analysis of data. 8) Test reports may be prepared.
| Personnel: | [6 empty] | 2 x Staff
| Planning Relationships: | [5 empty] | Close to production area. Close to administrative offices.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 15 |  |  |  |  | Height (mm): |  |  |  | 2700
| Notes: | [7 empty] | Storage of records: validation, routine testing , archived maintenance records. Communications: To all areas of the department and external areas. Separate data and voice outlets may be used where structure cabling solutions are not available.
| ADB | [7 empty] | Room Environmental Data
R1080
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1080 |  |  |  |  | Test room: equipment and data SSD No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | POS
| Supply Air: Final Filter Class | [12 empty] | >H10
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
| Type Of Control: | [13 empty] | N
| Daytime General Service Illuminance (Lux): | [13 empty] | 300
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | Y
| Colour Rendering Required Characteristics (Ra): | [13 empty] | 80
| Unified Glare Rating Limit (UGRL): | [13 empty] | 19
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
| Standby Lighting Grade - General Lighting: | [13 empty] | B
Standby Lighting Grade - Local Lighting:
| General Notes: | [5 empty] | Refer to SLL Lighting Guide 2 for more detailed guidance
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | High
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | R1080
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1080 |  |  |  |  | Test room: equipment and data SSD No room layout
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 1 doorset: 1 x personnel access (900mm)
| Windows: | [5 empty] | Not required
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1080
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1080 |  |  |  |  | Test room: equipment and data SSD No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | BAS101 |  | BASIN, medium, hospital pattern, vitreous china, no tap holes, no overflow, integral back outlet, 500W 400D. | [16 empty] | 1
| 3 |  |  | 3 |  |  |  |  | HOO024 |  | HOOK, hat and coat, 1 | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT002 |  | OUTLET cable, 13 amp | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT010 |  | SOCKET outlet, switched, 13 amp, twin | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT049 |  | CONNECTION UNIT, switched, 13 amp, flex outlet | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | OUT052 |  | CONNECTION UNIT, switched, 13 amp | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | OUT131 |  | SOCKET outlet data/voice, double. | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SNS757 |  | SINKTOP, single bowl, no tap holes, no overflow, right hand integral ribbed drainer, stainless steel, 1200W 650D. | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | STF275 |  | STORAGE UNIT, upper, cupboard, 2 door, 1 shelf, lockable, 550H 600W 300D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SWC025 |  | SWITCH, light | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | TAP359 |  | TAP, pillar, high neck, long lever, pair hot and cold, 1/2 in. | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | TAP892 |  | TAP, bib, 2x8 mm thermostatic mixer, automatic action, sensor operated, non-touch, | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | WOR032 |  | WORKBENCH, laboratory, 2000W 700D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | BOA013 |  | BOARD, display/notice, wall mounted, 900H 1200W | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS011 |  | DISPENSER, barrier cream, disposable single cartridge, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS013 |  | DISPENSER, paper towel, wall mounted | [16 empty] | 2
| 1 |  |  | 1 |  |  |  |  | DIS030 |  | DISPENSER, soap, disposable single cartridge, lever action, wall mounted | [16 empty] | 2
| 2 |  |  | 2 |  |  |  |  | CAB066 |  | CABINET, roller shutter, screen-hung/wall mounted, 430H 800W 430D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | CHA002 |  | CHAIR, height adjustable, medium back, swivel, 5 star base, on castors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | CHA063 |  | CHAIR, height adjustable, with arms, high back, swivel, 5 star base, on castors | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM031 |  | COMPUTER CPU | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM032 |  | COMPUTER VDT MONITOR | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | COM033 |  | COMPUTER KEYBOARD | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | DES027 |  | DESK UNIT, cantilever, cable management, adjustable legs, modesty panel, 1600W 800D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | DES028 |  | DESK RETURN UNIT, cantilever, cable management, adjustable legs, modesty panel, freestanding, 1200W 600D | [16 empty] | 3
| 1 |  |  | 1 |  |  |  |  | DRA056 |  | DRAWER UNIT, 2 drawer, lockable, on castors, 600H 410W 600D | [16 empty] | 3
| 2 |  |  | 2 |  |  |  |  | HOL006 |  | HOLDER, sack, with lid foot operated, medium, freestanding, 875H 430W 385D | [16 empty] | 3
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | R1080
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | R1080 |  |  |  |  | Test room: equipment and data SSD No room layout
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | REF060 |  | REFRIGERATOR, reagents, capacity 110 litre, external temperature gauge, automatic defrost, underbench, 850H 500W 590D | [16 empty] | 3
| ADB | [8 empty] | Room Data Sheet | [11 empty] | Y0706
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | Y0706 |  |  |  |  | Hold: contaminated returns, 4 SSD trolleys
| Room Number: | [16 empty] | Revision Date: | [5 empty] | 05/04/2013
| Activities: | [5 empty] | 1) Holding contaminated returns. 2) Holding locally generated clinical waste and associated waste products from the contaminated returns holding area and the wash room.
| Personnel: | [6 empty] | 2 x Staff Intermittent use
| Planning Relationships: | [5 empty] | Adjacent to the contaminated returns lobby. Within the wash room. Direct access to the general waste disposal room.
| Space Data: | [6 empty] | Area (m²): |  |  |  | 8 |  |  |  |  | Height (mm): |  |  |  | 2700
| Notes: | [7 empty] | Telephone outlet - option.
| ADB | [7 empty] | Room Environmental Data
Y0706
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | Y0706 |  |  |  |  | Hold: contaminated returns, 4 SSD trolleys
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| TEMPERATURE AND VENTILATION | [11 empty] | Requirements
Notes
| Permissible Space Temerature Range(dry bulb) (degC): | [13 empty] | 18 - 28
| Heating Design Temperature (dry bulb)(degC): | [13 empty] | 21
| Minimum Air Changes (AC/hr): | [13 empty] | 15
| Ventilation Type: | [13 empty] | S/E
| Pressure Relative to Adjoining Space: | [13 empty] | -ve
| Supply Air: Final Filter Class | [12 empty] | F9
| Permisiible Relative Humidity Range (%): | [13 empty] | Uncontrolled
General Notes:
LIGHTING
Type Of Control:
Daytime General Service Illuminance (Lux):
Daytime Specific Service Illuminance (Lux):
Nighttime General Service Illuminance (Lux):
Nighttime Specific Service Illuminance (Lux):
Local Task Illuminance (Lux):
| Colour Rendering Required: | [14 empty] | N
Colour Rendering Required Characteristics (Ra):
Unified Glare Rating Limit (UGRL):
| Emergency Escape Route Lighting Required: | [13 empty] | Y |  | In accordance with BS 5266 and Health Technical Memorandums
Standby Lighting Grade - General Lighting:
Standby Lighting Grade - Local Lighting:
General Notes:
RISK
Clinical Risk Category:
Non-clinical Business Continuity Risk Category:
General Notes:
NOISE
| Noise Intrusion (dB) 1hr day: | [13 empty] | 50
| Noise Intrusion (dB) 1hr night: | [13 empty] | -
| Noise Intrusion (dB) f night: | [13 empty] | -
| Maximum Internal Noise from M&E Services (NR): | [13 empty] | 40 |  | Total noise of MEP services under normal operation across the range 63Hz to 4kHz inclusive.
| Room Sound-insulation Parameters - Privacy: | [13 empty] | Not Private |  | Reference to Table 3 of the Department of Health 'Acoustics: Technical design manual 4032:0.6:England'.
| Room Sound-insulation Parameters - Noise Generation: | [13 empty] | Typical
| Noise Sensitiviy: | [13 empty] | Not Sensitive
Sound-insulation Rating (dB D nT,w):
General Notes:
SAFETY/FIRE
| Maximum Surface Temperature (DegC): | [13 empty] | 82
| Domestic Hot Water Discharge Temperature (DegC): | [13 empty] | 41/>55
| Maximum Cold Water Discharge Temperature (DegC): | [13 empty] | <20
General Notes:
| Type of Automatic Fire Detection: | [13 empty] | Smoke
General Notes:
| ADB | [7 empty] | Room Design Character | [12 empty] | Y0706
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | Y0706 |  |  |  |  | Hold: contaminated returns, 4 SSD trolleys
| Room Number: | [17 empty] | Revision Date: | [5 empty] | 05/04/2013
| Walls: | [5 empty] | Wall finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Wall finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Floor: | [5 empty] | Floor finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Floor finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
Ceiling:
Ceiling finishes to comply with Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England. Ceiling finishes to be selected using the "Selection Procedure for Finishes" included in 8941:06: England.
| Doorsets: | [5 empty] | Configuration, glazing, fire rating, security, etc. to be determined by Project Team. HTM 58 (Mar-2005) Refer to HBN 00-04 (May-2007) for effective clear door widths. 1 doorset: 1 x personnel, bed, trolley, wheelchair & equipment access (1700mm); Through vision, self-closing delayed action, automated/interlocked system; trolley access. The need for a door depends on the plan adopted.
| Windows: | [5 empty] | Non-essential - Project Option If provided must be sealed Designation to be validated against current documentation (HTM 55 archived).
Internal Glazing:
Not required
| Hatch: | [5 empty] | Not required
Notes:
All finishes selected must have an appropriate risk assessment to accompany the design decision. Infection Control must be consulted as described in Performance Requirements for Building Elements Used in Healthcare Facilities 8941:0.6 England.
| ADB |  |  |  |  | Schedule of Components by Room | [15 empty] | Y0706
| Project: | [5 empty] | ADB2013 |  |  |  |  | Activity Database version 2013 © Crown Copyright
| Department: | [5 empty] | HBN13 |  |  |  |  | Sterile Services Department
| Room: | [5 empty] | Y0706 |  |  |  |  | Hold: contaminated returns, 4 SSD trolleys
| Room Number: | [18 empty] | Revision Date: | [5 empty] | 05/04/2013
|  | Quantity | [20 empty] | Alt. Code | [5 empty] | Grp
New |  |  | Trans |  | Total |  |  |  | Code |  |  | Description
| 1 |  |  | 1 |  |  |  |  | OUT005 |  | SOCKET outlet, switched, 13 amp, single | [16 empty] | 1
| 2 |  |  | 2 |  |  |  |  | SHE222 |  | SHELF, 1200W 400D | [16 empty] | 1
| 1 |  |  | 1 |  |  |  |  | SWC025 |  | SWITCH, light | [16 empty] | 1
Activity DataBase
2016-12-29